## Sheet1
Open HFNC | Bill of Materials
System | Subsystem | Part | Quantity | Material | Manufactured or COTS | Description | Reference Draft | COTS - National (Chile) | COTS - International
Heating-Box | Heating plate box | Box base | 1 |  | Manufactured | Base of the box that contains the parts that allow the heating of the heating plate and the beaker glass.
|  | Box top | 1 |  | Manufactured | Cover of the box that contains the parts as descripted above.
|  | M4 rivet nut | 14 |  | COTS | Rivet nut on box base used for linking it with the cover trough the DIN912 M4x10 Screw |  |  | https://www.mcmaster.com/95105A167/ (PENDIENTE, NO ES LA MISMA)
|  | Heating plate | 1 |  | Manufactured | Cirucular plate that receives the heating resistance, distributing the heat, mainly in one of his faces where the beaker glass stands..
|  | Oring 2-343 | 1 |  | COTS | Seal that isolates thermal conduction from the heating plate to the top box. Also functions as water leak seal. |  | Parker O-Ring Catalog
|  | Heater fixing bar | 1 |  | Manufactured | Aluminium bar (125x12x1). Functions as supporting link between the heating plate and the top box.
|  | DIN912 M4x40 | 2 |  | COTS | DIN912 M4x0.7x40-8.8 socket head screw Holds heater fixing bar to box top |  |  | https://www.mcmaster.com/91290A184/
|  | DIN 934 M4 nut | 2 |  | COTS | DIN912 M4x0.7 Nut - Holds heater fixing bar to box top |  |  | https://www.mcmaster.com/90592A090/
|  | DIN 912 M4x10 | 10 |  | COTS | DIN912 M4x0.7x10-8.8 socket head screw - Secures Box base to box top fitting itself in to M4 Rivet Nut |  |  | https://www.mcmaster.com/91290A144/
| Electric parts | Thermo couple | 1 |  | COTS | Thermocouple Type K M6 Thread. Measures temperature from Heating plate, REXC100 controller reads it. |  |  | https://www.mpja.com/Thermocouple-Type-K-M6-Threaded-Tip/productinfo/32106%20MI/
|  | Heater 630W | 1 |  | COTS | Cartridge heater, 220V - 630W. 5/8' DIamater 3' Long. Converts electrical energy in to heat |  | Armando Mendez <amendez@fabrestel.cl> | https://www.mcmaster.com/35025K473/ (NOTA: IDEALMENTE 630W)
|  | REXC100 controller | 1 |  | COTS | PID temperature controller.Adjustable Setpoint 0 - 400ºC |  |  | https://www.mcmaster.com/38615K73/ https://www.amazon.com/-/es/REX-C100-48-100-controlador-temperatura/dp/B07KXPLZTQ
|  | SSR-25DA | 1 |  | COTS | Solid State Relay; controls current flow to Heater 630W (opening it or closing it) in response to the REXC100 signal |  |  | https://www.amazon.com/Artshu-M%C3%B3dulo-SSR-25DA-entrada-24-380VAC/dp/B07JZF6973/
|  | On-Off Switch | 1 |  | COTS | On/Off switch for electrical energy, from OpenHFNC team. |  |  | https://www.mcmaster.com/7030K45/
|  | IEC C14 male panel connector | 1 |  | COTS | C14 male power connector |  | https://altronics.cl/conector-poder-panel-macho-c14?tag=C14 | https://www.amazon.com/Baomain-Panel-Power-Socket-Connector/dp/B00WFYS1HS
|  | 220 AC power cable | 1 |  | COTS | 220 AC power cable- C14 female power connector |  |  | https://www.mcmaster.com/power-cords/voltage~250v-ac/
|  | DIN 913 M5x5 | 2 |  | COTS | DIN 913-M5x5 socket screw flat end-Sets heating cartridge to heating box |  |  | https://www.ametric.com/product/85223/m5x5-din-913
|  | DIN 7991 M3x10 | 2 |  | COTS | DIN 7991 M3x0.5x10-8.8 countersunk head cap screw.Sets panel connector C14 to top box |  |  | https://www.mcmaster.com/91294A130/
|  | DIN 934 M3 nut | 2 |  | COTS | DIN934 M3x0.5 nut. Sets Panel C14 power connector to top box. |  |  | https://www.mcmaster.com/90592A085/
| Glass vase & flow meter fixer | Glass fixer plate | 1 |  | Manufactured | Plaque that is set to heating plate and allows the beaker glass and the flowmeter to stay in place. |  |  | -
|  | DIN 912 M4x10 | 2 |  | COTS | DIN912 M4x0.7x10-8.8 socket head screw - Sets the glass fixer plate to the heating box. |  |  | https://www.mcmaster.com/91290A144/
|  | Oring 2-159 | 1 |  | COTS | Oring that functions as elastic band between beaker glass and the glass fixer plate. |  | Parker O-Ring Catalog
| IV stand fixer | IV stand fixer plate | 2 |  | Manufactured | Plaque that is set to the heating box and allows in his other end to be fixed to the IV stand.
|  | Star knob 1/4 UNC | 1 |  | COTS | Star Knob 1/4-20UNC. For stand fixer plate to IV stand fixation |  |  | https://www.amazon.com/Big-Horn-19707B-5PK-Star-Thru/dp/B01DO22H5Q/ref=sr_1_3?_encoding=UTF8&c=ts&dchild=1&keywords=Knobs+%26+Hand+Wheels&qid=1596129827&s=industrial&sr=1-3&ts_id=16412791
|  | DIN 912 M4x10 | 6 |  | COTS | DIN912 M4x0.7x10-8.8 socket head screw - (4) are for setting the stand fixer plate to heating box. (2) are for attaching both stand fixer plates. |  |  | https://www.mcmaster.com/91290A144/
|  | DIN 934 M4 nut | 2 |  | COTS | DIN 934 M4x0.7 Nut.Used for attaching both stand fixer plates. |  |  | https://www.mcmaster.com/91415A020/
|  | Split washer DIN 124B D4 | 2 |  | COTS | DIN 124B Split Washer Inner diameter 4mm.Used for attaching both stand fixer plates.
Ventilation pipeline | Flow pipelines | Flowmeter | 1 |  | COTS | Flujometro 10-100 LPM (0.6-6 m3N/h) - Conections G1/4 BSP female |  |  | https://www.grainger.com/product/BROOKS-10-to-100-lpm-Variable-Area-16X857
|  | Corrugated tube | 1 |  | COTS | Tube that carries air/oxygen flow from flowmeter to the beaker glass. |  | http://www.megamed.cl/product/tubo-corrugado
|  | Heated tube | 1 |  | COTS | Heated tube that carries air/oxygen humidified flow from beaker glass to transfer tube. |  |  | https://www.directhomemedical.com/HBT300-hybernite-heated-cpap-hose-tube.html
|  | Transfer tube | 1 |  | COTS | Tube that carries air/oxygen flow from heated tube to nasal cannula. .Also serves as flow area smooth reduction tube between the components at his ends. |  |  | https://www.amazon.com/Hand-Drum-Pump-Siphon-Polyethylene/dp/B00ZZY2EFU Dukshin, Siphon Hand Pump (Product Number: DP-14)
|  | BSP 1/4 - Hose 7/8 | TBD |  | COTS | Fitting Connection between flowmeter output and corrugated tube
|  | BSP 1/4 - Hose 1/4 | TBD |  | COTS | Fitting Connection between 1/4' high flow medical tube and flowmeter input |  |  | https://www.mcmaster.com/1075T22/
|  | Y conection | TBD |  | COTS | Y connector that receives air and oxygen and delivers to both 1/4' high flow medical tubes.
|  |  | TBD |  | COTS
| Beaker glass - Top connectors | Beaker glass 600ml, diam: 90mm | 1 |  | COTS | Beaker Glass 600ml diameter 90mm. For distilled water. Receives heat coming from the heating plate.Where water evaporation for air/oxygen flow humidification occurs. |  |  | https://www.duran-group.com/en/produkte-loesungen/laborglas/produkte/kochglaeser-allgemeines-laborglas/becher.html - Cat. No.: 21 106 48
|  | Top lid body | 1 |  | Manufactured | Beaker glass top lid body. Seals the gaps and create connections for the heating tube (oulet flow) and corrugated tube (inlet flow)
|  | Heated tube-top lid connector | 1 |  | Manufactured | Connector between heating tube and top lid body.
|  | Corrugated tube-top lid connector | 1 |  | Manufactured | Connector between corrugated tube and top lid body
|  | Paddle | 1 |  | Manufactured | Paddle that provides separation inside the beaker glass and allows the entering flow to moisturize by forcing it to pass near water in evaporation process .
|  | DIN 912 M4x16 | 1 |  | COTS | DIN912 M4x0.7x40-8.8 socket head screw -Sets Paddle to Top Lid Body |  |  | https://www.mcmaster.com/91290A154/
|  | Oring 2-007 | 1 |  | COTS | Seal between paddle and top lid body |  | Parker O-Ring Catalog
|  | Oring 2-018 | 1 |  | COTS | Seal between corrugated tube amd top lid connector |  | Parker O-Ring Catalog
|  | Oring 2-019 | 1 |  | COTS | Seal between corrugated tube-top lid connector and top lid body |  | Parker O-Ring Catalog
|  | Oring 2-022 | 1 |  | COTS | Seal between heated tube-top lid connector and heated tube |  | Parker O-Ring Catalog
|  | Oring 2-026 | 1 |  | COTS | Seal between heated tube-top lid connector and top lid body |  | Parker O-Ring Catalog
|  | Oring 2-151 | 1 |  | COTS | Seal between top lid body and beaker glass |  | Parker O-Ring Catalog
| Nasal cannula | Nasal cannula body | 1 |  | Manufactured | BIocompatible 3D printed nasal cannula Delivers humidified air/oxygen flow to the patient through his nostrils.
|  | Oring 2-014 | 1 |  | COTS | Seal between nasal cannula body and transfer tube |  | Parker O-Ring Catalog
|  | Elastic band | hay que crearlo |  | COTS | Elastic band used to hold nasal cannula to the patients head. |  | https://www.amazon.com/dp/B07K7Y62BR/ref=sspa_dk_detail_4?psc=1&pd_rd_i=B07K7Y62BR&pd_rd_w=8aZPO&pf_rd_p=48d372c1-f7e1-4b8b-9d02-4bd86f5158c5&pd_rd_wg=pa3t0&pf_rd_r=CVF28XY01B6MS6VMGXGR&pd_rd_r=e60ab4c0-170c-448d-a493-7acf86a5691c&spLa=ZW5jcnlwdGVkUXVhbGlmaWVyPUFIWVZNMlAxTUgzWkomZW5jcnlwdGVkSWQ9QTAwNjAzNzQyRUVDWVdRNEJRODVKJmVuY3J5cHRlZEFkSWQ9QTA4Mjc5NjVTVVpGNjZJMlgzSk4md2lkZ2V0TmFtZT1zcF9kZXRhaWwmYWN0aW9uPWNsaWNrUmVkaXJlY3QmZG9Ob3RMb2dDbGljaz10cnVl